## Specifications
Bid Ref. B26-19 High Energy Linear Accelerator System Date:17/09/2026
Required Specifications for: | High Energy Linear Accelerator System
Brand | To be specified
Model | To be specified
Made | To be specified
FDA Cleared | Yes
Delivery | To be specified
Core System & Beam Characteristics
Photon Energies | Dual photon energies: 6 MV and 15 MV. Must include High-Dose-Rate / Flattening Filter Free (FFF) modes for rapid delivery, and capability to deliver total body irradiation.
Electron Energies | Discrete energies (e.g., 6, 9, 12, 15, 18, and 21 MeV) for superficial and semi-deep tissue treatments, and capability to deliver total skin electron therapy.
Collimation System | Integrated Multi-Leaf Collimator (MLC) with high spatial resolution (leaf width ≤ 5 mm at isocenter over central field, or micro-MLC equivalent
Dose Rate Output | Variable high-dose rate for FFF photon modes to minimize intra-fraction motion.
Isocenter Precision | Mechanical and radiation isocentric accuracy ≤ 0.5 mm sphere of uncertainty for Stereotactic Radiosurgery (SRS) and SBRT. Non-mechanical wedging capability for static fields.
Imaging & Motion Management
On-Board Imaging (OBI) | High-resolution 3D Cone-Beam CT (CBCT) and 2D kilovoltage (kV) / megavoltage (MV) planar imaging with automated registration algorithms.
Respiratory Gating & Surface Tracking | Integrated Optical Surface Management System (SGRT / OSMS) or 4D-CT respiratory gating capability for real-time motion tracking (e.g., Deep Inspiration Breath Hold).
Patient Positioning Couch: | Fully integrated 6-Degrees-of-Freedom (6DoF) robotic couch allowing remote translational, pitch, roll, and yaw corrections and capability to apply couch shifts/rotations from outside bunker.
Software & Infrastructure Integration
Treatment Planning System (TPS) | Minimum 2 concurrent user licenses with Monte Carlo (or equivalent) or Collapsed Cone Convolution calculation algorithms for advanced IMRT, VMAT, SRS, and SBRT planning.
Oncology Information System (OIS) | Full bidirectional compatibility with KMC's Hospital Information System (HIS) and PACS utilizing standard DICOM, DICOM-RT, and HL7 interfaces.
Dosimetry & Quality Assurance | Complete QA hardware and software suite including 3D motorized water phantom, 2D matrix array detector for patient-specific plan verification, and daily QA equipment or cylindrical detector tool for patient-specific VMAT plan verification.
Portal Dosimetry | Capability of performing Portal Dosimetry to verify VMAT dose
Site Preparation & Infrastructure Resilience
Power Conditioning & UPS | Bidder must include a dedicated, medical-grade online double-conversion UPS system capable of supporting 100% equipment load during active treatment, beam-on sequences, and image processing for a minimum of 15–20 minutes, including transient surge protection for generator-grid switches.
Bunker Audit & Adaptation | Bidder shall perform an architectural and radiation shielding audit of KMC's existing treatment bunker, providing comprehensive HVAC, chilled water, electrical, structural, and lead/concrete shielding specifications.
Cooling Systems | Dual redundant external chiller units dedicated to maintaining optimal operating temperatures for the magnetron/klystron wave-guide cooling loop.
Financial & Acquisition Models
Option A Under the CAPEX model, KMC acquires full ownership of the LINAC system upon clinical hand-over. | • 1. Equipment Price: Itemized cost breakdown covering Hardware, Software, Initial Dosimetry Suite, Shipping/Insurance, and Site Preparation. • 2. Milestone Payments: 15% upon Contract Signing; 55% upon Equipment Delivery to KMC; 20% upon Successful Installation & Acceptance Testing; 10% upon First Patient Treated (or 90 days post-handover).
Warranty | . Warranty & CAMC: 5-years all-inclusive warranty (parts, labor, glassware, software updates). Post-warranty Comprehensive Annual Maintenance Contract (CAMC) locked for Years 3–10 with guaranteed uptime ≥ 98% and on-site engineer response ≤ 24 hours.
Option B Revenue Sharing / Operational Model (OPEX), Under the OPEX model, the vendor retains asset owenrship and partners with KMCto share operational success. | • 1. Capital Structure: Zero upfront capital outlay for equipment from KMC. The vendor finances, delivers, installs, and maintains the full LINAC system. • 2. Fee Structure: Fixed fee paid to vendor per treated fractions. • 3. Tiered Volume Discount: reduced per-fraction charges once agreed monthly volume thresholds (N fractions) are surpassed. • 4. Term Duration: Long-term operational alignment for a period of 10 years. • 5. Responsibilities: Vendor covers all maintenance, replacement parts (including magnetron, klystron, and electron gun), software upgrades, and uptime guarantees (≥ 98%). KMC provides clinical staffing, operational management, billing, and facility utilities. • 6. End-of-Term Options: At contract expiration, KMC retains options to purchase the equipment at Fair Market Value (FMV), extend the agreement, or request vendor-funded hardware replacement.
The vendor must guarantee comprehensive site engineering support and power isolation systems
The vendors must guarantee the supply, installation, testing, commissioning, and long-term technical support of an advanced, high-energy Linear Accelerator (LINAC) system. The system must feature fully integrated Image-Guided Radiation Therapy (IGRT), Volumetric Modulated Arc Therapy (VMAT), Stereotactic Body Radiation Therapy (SBRT), and Stereotactic radiosurgery (SRT/SRS) capabilities