## Specifications
Bid Ref. B26-24 Contrast Injector for Cath Lab Date:16/10/2026
Required Specifications for: | Contrast Injector | Supplier's Answer
Brand | To be specified
Model | To be specified
Made | To be specified
FDA Cleared | To be specified
Delivery | To be specified
General
Application | Suitable for Angiography, Coronary Angiography, Interventional Procedures and TAVI / Structural Heart Procedures
Configuration | Single-head / Single-syringe high-pressure injector
Injection System | Electromechanical angiographic contrast injector
INJECTION PERFORMANCE
Maximum Programmable Pressure | ≥ 1,200 psi
Maximum Flow Rate | ≥ 50 mL/sec
Minimum Flow Rate | ≤ 0.1 mL/sec
Flow Rate Adjustment | Fine programmable adjustment throughout the operating range
Injection Volume | Programmable from approximately 1 mL up to syringe capacity
Volume Increment | ≤ 0.1 mL preferred
Injection Phases | Minimum 4 programmable phases
Rise / Fall Time | Programmable, approximately 0–10 seconds or better
Injection Delay | Programmable delay, 0–99 seconds or better
SYRINGE
Syringe Capacity | Minimum 150 mL
Syringe Compatibility | 150 mL and preferably 200 mL
Syringe Loading | Easy and rapid loading/unloading with safety mechanism
Syringe Recognition | Automatic or configurable syringe size recognition
PRESSURE & SAFETY
Pressure Monitoring | Real-time pressure monitoring
Pressure Limit | Programmable and adjustable pressure limit
Over-Pressure Protection | Automatic alarm and protection against excessive pressure
Air Management | Air detection/management system
Safety System | Visual and audible alarms with appropriate safety interlocks
DISPLAY & CONTROL
Display | Minimum 12-inch color touchscreen or equivalent advanced graphical interface
Display Information | Flow rate, volume, pressure, injection time, injection phase, delay and syringe status
User Interface | Simple and user-friendly graphical interface
Protocol Memory | Ability to save and recall multiple injection protocols
Remote Control | Remote operation capability preferred
Start / Stop Control | Hand switch, foot switch or equivalent
CONTRAST MANAGEMENT
Contrast Warming | Integrated contrast warming system
Warming Temperature | Approximately 37°C
Temperature Safety | Temperature monitoring and appropriate safety protection
CATHLAB / TAVI INTEGRATION
X-Ray Interface | Interface with Cathlab X-ray / imaging system
Imaging Synchronization | Capability to synchronize contrast injection with image acquisition
TAVI Application | Suitable for aortic root, peripheral and other angiographic injections required during TAVI
Cathlab Compatibility | Compatible with standard Cathlab angiography systems
INSTALLATION & SERVICE
Installation | Delivery, installation, commissioning and functional testing included
Training | User training for Cathlab staff included
Technical Support | Authorized local technical/service support
Warranty | Minimum 2 years
Preventive Maintenance | Preventive maintenance according to manufacturer's recommendations
Spare Parts | Availability of spare parts and consumables through authorized supplier
Regulatory Compliance | CE and/or FDA clearance, as applicable
Documentation | User manual, technical specifications, service/maintenance documentation and regulatory certificates